[image: image1.jpg]ULTRASOUND

DIAGNOSTIC SONOGRAPHY AND THERAPEUTIC
APPLICATION OF ULTRASOUND

ki FE |






[image: image19.jpg]



[image: image20.jpg]



[image: image21.jpg]



[image: image22.jpg]


Ultrasound, that is mechanical oscillation of sufficiently high frequency is applied in medicine for both diagnostic and therapeutic purposes. Diagnostic ultrasonic applications (ultrasonography, medical sonography) are non-invasive methods, that show enormous development in the last decades (2D, 3D, 4D, Doppler methods, etc.). Biomedical therapeutic applications include mostly those of rheumatology, for the treatment of musculoskeletal disorders (especially in arthrosis), but ultrasonic scaler, that is used for the removal of the dental calculus (tartar) applies frequencies above 20 000 Hz as well.

REVIEW OF THE THEORY

[image: image23.jpg]


PROPAGATION, DAMPING AND REFLECTION OF ULTRASOUND

Ultrasound is a mechanical wave, that is longitudinal in liquids and gases, and it is both longitudinal and transverse in solids.

Propagation of ultrasound

[image: image24.jpg]



Ultrasound propagates with velocity (c), that is characteristic for the medium, and behaves as waves (general phenomena: reflection, refraction).

Damping of ultrasound
Part of the energy, that propagates in the ultrasound is absorbed in the medium, in the form of heat. According to the general attenuation law of radiation:


[image: image2.wmf]x

e

J

J

×

-

=

m

0

,



(1)
where J and J0 are the intensities measured at distance x and 0, respectively, and μ is the absorption coefficient.

[image: image25.jpg]


[image: image26.jpg]


The damping is used to characterize the decrease of the intensity, which is similar to the concept of optical or X-ray density. Damping is defined as: 
[image: image3.wmf](

)

e

x

J

J

lg

10

dB

lg

10

0

m

a

=

=

. As μ is proportional to the frequency (f ), this holds for α as well. Absorption property of tissues is often characterized by specific damping , that has a following unit: (dB/(cm·MHz)).

Reflection of ultrasound

At the boundary surface of different media part of the incident ultrasound wave is reflected (reflection) and part penetrates into the medium (Fig 1). If the direction of propagation of the incident ultrasound wave is not perpendicular to the boundary surface (skew incidence), the ultrasound wave will change its direction of propagation (refraction), similarly to the refraction of light.

Acoustic impedance (Z ) is an important parameter that characterizes acoustic properties of the medium. It is defined as:


[image: image4.wmf]c

Z

×

=

r

,



(2)
[image: image27.jpg]


where ρ is the density of the medium, and c is the velocity of the ultrasound in the medium. The measure of reflection from a surface is the reflexivity (R), which is defined as the ratio of the reflected (JR) and the incident (J0) intensity. 


[image: image5.wmf]2

2

1

2

1

0

R

÷

÷

ø

ö

ç

ç

è

æ

+

-

=

=

Z

Z

Z

Z

J

J

R

,



(3)
Reflexivity can be calculated from the formula above, using the known acoustic impedances (Z1, Z2) of the media.

THE SOURCE AND THE DETECTION OF ULTRASOUND

[image: image28.jpg]


Direct piezoelectric effect

Certain solid materials (e.g. quartz, BaTiO3) accumulate electric charge and get polarized when mechanical strain (pressure, tension) is applied on them. This results in measurable potential difference between the electrodes placed on the opposite sides of the crystal (Fig 2. a). Detection of mechanical oscillations as of ultrasound (that means converting ultrasound into electric signal) is based on this phenomenon.

[image: image6.jpg]\ tension

coincident separated separated
centers+ centers I centers +
of charges of charges of charges





Fig 2. a. Direct piezoelectric effect.
Inverse piezoelectric effect

Piezoelectric effect is reversible. If an electric voltage is applied on the electrodes of the piezoelectric material, it becomes deformed (shrank or stretched), thickness of the crystal changes (Fig 2. b). Sources of the ultrasound are constructed based on this inverse piezoelectric effect.

[image: image7.jpg]expansion
shrinking T P+ + + + +

coincident separated + separated
+ centers 7 centers * centers
of charges of charges ~ of charges





Fig 2. b. Inverse piezoelectric effect.

[image: image29.jpg]



An almost instant emission and detection of ultrasound is essential in medical ultrasonography.

Ultrasound source

[image: image30.jpg]


Alternating voltage of high frequency, that is converted into ultrasound is produced by a sine wave oscillator (see chapter 18. SINE WAVE OSCILLATOR). This electric signal is actually a short wave package of several MHz frequency, that contains only a few number of periods of a sine wave, thus in fact, it is an electric pulse (Fig 3). This electric pulse is conducted to the transducer (radiating and detector head), which contains the piezoelectric material equipped with the electrodes.

Transducer converts the electric energy into the ultrasound (inverse piezoelectric effect), that is directed into the examined part of the body. As it was mentioned, this ultrasound pulse (“transmission”) is short, lasts couple of time periods (wave packet), just like the pulse-like ultrasound whistles of the bat.

Detection of ultrasound

[image: image31.jpg]


In the “silence”, that follows  the transmission of the ultrasound pulse the transducer “switches into receiver mode”, and waits for the echo signals reflected from the tissue surfaces of the examined body part. After certain time the echo of reduced intensity reaches and deforms the transducer material (just like it happens in the ultrasound-sensitive ear of the bat). As a consequence of the electric polarization in the transducer, the ultrasound is converted back into the electric oscillation (direct piezoelectric effect), and this signal is then conducted through a cable into the electronic amplifier (see chapter 17. AMPLIFIER).
[image: image32.jpg]velocity
color code

1.8MHz
WF 300Hz

color/coded selected
velocit direction




APPLICATION OF THE ULTRASOUND: MEDICAL ULTRASONOGRAPHY, DIAGNOSTICS
In the last decades, large progress was made in the diagnostic applications of ultrasound. Method is based on the reflection of the ultrasound from the surfaces of biological tissues having different acoustic impedances. Distance of the examined boundary surfaces is calculated from the measured time between the emission of ultrasound pulse and the arrival of the echo signal, and the image of these surfaces is reconstructed.
As the frequency of the reflected ultrasound changes if the reflecting surface is moving (Doppler effect, see later), the frequency spectrum of the echo contains functional information as well. 

Pulse-echo principle
The distance (d)  between the reflecting boundary surface and the transducer can be calculated from the time (Δt) elapsed between the emission of the wave-packet-like ultrasound pulse into the medium and the time of the arrival of the ultrasound echo signal, if the velocity of the sound (c) is known in the given medium (Fig 3), as follows:

 
[image: image8.wmf]2

t

c

d

D

×

=

.



(4)
[image: image33.wmf](

)

dB

lg

10

0

J

J

=

a

[image: image9.jpg]SOURCE

emission of the
electric pulse

DETECTION

elecric

t

t+At

echo pulse V\/W\/V

piezo-
transducer

skin

A

body

A~

poundary surfag,

emitted
US pulse

d

y

poundary surfge,

<

reflected
US pulse





Fig 3. Pulse-echo principle.
The time delay and the amplitude decrease of the echo signal compared to the emission pulse contains all the information about the location of the organs examined in the body (Fig 4). Wave-packet release time (transmission) and detection time (reception) changes periodically in the transducer. The main frequency of the applied ultrasound is around several MHz, thus in the emitted wave-packet there are only a few number of sine wave periods at 1 μs time (the approximate duration of one pulse). The repetition frequency of the emission-detection cycle is around 1 kHz (Fig 5).
[image: image10.jpg]RANS- TRANS- TRANS-
MISSION RECEPTION MISSION RECEPTION MISSION RECEPTION

\

2"echo techo  2"™echo

2"echo 1*'echo

1*echo

approx.

1us 1ms




Fig 5. Time sequence of the emission and detection of the ultrasound.

[image: image34.wmf]2

2

1

2

1

0

R

÷

÷

ø

ö

ç

ç

è

æ

+

-

=

=

Z

Z

Z

Z

J

J

R

Diagnostic ultrasound imaging methods
One-dimensional A-image (amplitude modulation): Single transducer produces one ultrasound beam that propagates in a straight line. In the image, the horizontal axis represents time or the corresponding distance, and the vertical axis denotes the intensity of the reflected echo signal (Fig 6., center).

One dimensional B-image (brightness modulation): single transducer produces one ultrasound beam that propagates in a straight line. The brightness of pixels on the horizontal time scale is proportional to the amplitude of the reflected ultrasound signal. Brightness is encoded usually on a gray scale (Fig 6., bottom, brightness changes along the blue line).
Two dimensional B-image (2D, brightness modulated): an array of several transducers are specially synchronized the way, that the direction of the wave front can be changed, and a plane (fan shaped) is scanned. A series of one dimensional B-images at different angles appears on the screen (curved array, Fig 7., on the right). Real image is reconstructed of one section of the scanned region of the biological sample. 

[image: image35.wmf]2

Δ

t

c

d

×

=

M-image (motion): In the 2D B-image one direction is chosen, and a sequence of one-dimensional B-images is plotted vertically, as a function of time (horizontal time scale) (Fig 7., on the left). We get the information on the position of the given surface as a function of time.
Reconstructed 3D-image (tomography): Three dimensional image of the biological object can be reconstructed from the 2D B-images of a large number of parallel scanned planes (Fig 8.), that can be processed  further by the computer, thus it can be turned at any directions or different sections can be calculated. 

Reconstructed 4D-image: Sequence of several reconstructed 3D images, where time is the 4th dimension, and an ultrasound movie is created.

Doppler methods
The Doppler effect:
[image: image36.jpg]


The Doppler effect is a frequency (and wavelength) change of the wave as a result of the relative motion of the source and the observer (Fig 9).

[image: image11.jpg]



[image: image37.jpg]UT COLLEGE OF
RINAR

d (depth)

<ID:Q¥
C

Y MEDICINE<CH

YPURZ=16 02u03 . 93

selected direction



Fig 9. Doppler-effect: sound of the approaching source appears of higher pitch (shorter wavelength), while sound of a source moving away appears lower-pitch (longer wavelength) for the observer.

Similarly to the phenomenon described above, the frequency of the ultrasound that was reflected from a moving boundary surface (considered as a sound source) differs from the original frequency as follows:


[image: image12.wmf]÷

ø

ö

ç

è

æ

±

@

c

f

f

u

2

1

0

,




(5)
where f and f0 is the frequency of the reflected and the incident ultrasound wave respectively, c is the velocity of the ultrasound in the medium and  is the velocity component of the reflecting surface in the direction of the ultrasound propagation. The positive sign is used in case the source and the observer are moving towards each other, while the negative sign is applied in case the source and the observer are moving away from each other. Frequency of the ultrasound that is reflected from a still surface does not change. The Doppler shift (f – f0) is proportional to the relative velocity ( /c) and to the incident frequency. The latter is set the way, that the Doppler shift falls in the audible frequency range (several kHz), thus it can be observed directly using a loudspeaker.

[image: image38.jpg]SﬂUND-WAVEV PORTRAIT IN THE FLESH

arlke dovice prod



Doppler time-velocity image: At the selected section of the indicated direction in the 2D B-image, Doppler-frequency-shift that corresponds to the velocity of the observed surfaces is plotted as a function of time (Fig 10, right).
Color coded Doppler image: the 2D B-image (scaled gray) of the biological object and the color coded velocity information of the moving parts are shown simultaneously, at the same image (Fig 10, left).
Doppler flow meter: makes possible to measure the velocity of the blood flow in larger blood vessels:


[image: image13.wmf](

)

Q

u

cos

f

f

f

c

0

0

2

-

=

,




(6)
Where  is the velocity of the blood flow, c is the ultrasound velocity in the medium, f – f0 is the Doppler shift and Θ is the angle between the ultrasound beam and the axis of the blood flow (Fig 11.)

THE DESCRIPTION OF THE EQUIPMENT

Fig 12. shows the front panel and the main functions of the ultrasound  diagnostic system HITACHI EUB-40 (on the cover).

[image: image14.jpg]Doppler-cursor Doppler Doppler Doppler
. (sample length) signal magnification gain mode (in) Doppler
main power sound volume

3
© Power

freeze the
Doppler-image

start scan
o . N store data
Glgldlalal's] cl'al sfgfe] ] o o 1 R e
o e lw ) Y = |28
keyboard for . D MDY Wt - ESEOL = |5 > US power
. HI0G0DDDDDDBBE
typing text - s f
. & '3 Ix‘ o I-v I'e I‘“ I\M I, -I\ ,I’ 5] 3 3 q ocus
Image [ e Q B.BM. M
. . . HARACTER . SPACE ) ¥o STORE, == A § 5
orientation (flip) ‘ image selection
= aurg D IV O | O [« @ aQ 2
. S os i § e { P e femi W b\
gain ® o 1% . [%cn T depth gain control
(auto/manual) sl [ LS o\l o %o ‘

o V3 e . ! | US-transducers
body mark e oRs @} [H : 1- 3.5 MHz
selection = | A i (2_ 5 MHz)‘

| A
angle of j# [
cursor |
adjustment |
) d )
¥ AH s
{ window
distance — N_ (single/dual)
measurement cursor zoom US-transducer e AeivE
start/stop in/out selection ) the i Wwindow
area cursor (1- 3.5MHz gain € image
measurement

(x-y positioning) 2- 5MHz) (general) selction




Fig 12. The front panel of the ultrasound diagnostic system HITACHI EUB-40

PLAN OF THE EXPERIMENT

DEMONSTRATION

Detection of the blood flow based on the Doppler-effect

The ultrasound transducer (probe) producing 8 MHz frequency ultrasound is placed above the artery of the lower arm or the neck in proper angle. The ultrasound, reflected from the formed elements of the blood is converted to the electric signal in the second transducer (detector). The difference between the frequency of the original and the reflected ultrasound signal is proportional to the flow velocity of blood (Doppler effect). After amplification, this difference signal is audible e.g. by a loudspeaker. 

The radar principle – distance measurement by ultrasound – ultrasound A-image

Using a special ultrasound-head, short ultrasound pulses are emitted in the air towards a reflecting surface, and the same ultrasound head detects the echo signal. Both the signal of the emitted ultrasound and the reflected echo signal are visualized on the screen of the oscilloscope. 
[image: image39.jpg]cross section of the Tingers ot the rubber gloves
tank (filled with different media)
/

\

_rubber membrane

transducer

O
OOO&




[image: image15.jpg]reflecting
surface
us
At transducer

A-image
(emission)





Fig 13. Distance measurement using ultrasound in the air.

[image: image40.jpg]d (depth), t (time)

v

)
°
3
£
a
3
©

one dimensional B-image along the axis of the eye




Measure the time between the two signals (Δ t ), and using the known sound velocity in the air (c = 340 m/s), calculate the distance of the reflecting surface from the transducer (d = c·Δ t /2). Measure the transducer-to-reflecting surface distance by a measuring tape as well and compare it with the calculated value.

Measurements using a water-filled phantom (model)

1. Measure the distances of the individual threads of a fishing line thread system from the transducer on the screen (use the cursor). Make a calibration diagram, that means, plot the distance data as a function of the real thread-to-transducer distances.
2. Measure the depth of penetration of the different transducers (3.5 MHz, 5 MHz) by moving the stick inside the phantom.

3. Measure the lateral and axial resolution by changing the position of the mobile fishing line threads, at different distances from the transducer.

4. Place the water-filled rubber gloves into the phantom (Fig 14.), and observe the images of different sections of different layers.

5. Place the air-filled rubber gloves into the phantom, and note the shadow formation.

6. Place the glycerol-filled and the alcohol-filled rubber gloves into the phantom. Draw the observed shape of the central finger, note the deformations of the shape (Fig 15). Explain the phenomenon. Measure the cross-section in every case for different media, and estimate the change of the velocity of sound as follows:


[image: image16.wmf]medium

water

water

medium

d

d

c

c

=

.


(7)

7. Turn the device into BM mode and observe the M-image of the ruler that is moved back and forth in the water of the phantom. Turn the device into Doppler mode and observe the velocity changes of the ruler movement in time (listen to the sound corresponding to this movement by loudspeakers).

Measurements on the human body

1. Observe the image of the fingers of a volunteer student, who places slowly her/his open hand into the water of the phantom.

2. The arterial sonography of the carotid artery (arteria carotis) of a volunteer student. After freezing the image at a proper moment measure the inner cross-section of the artery, its distance from the skin surface, and its area (hand-draw the contour of the artery by cursor to highlight the area).

3. Turn on the BM-mode and observe the time dependent pulsation of the artery.

4. Switch on the Doppler mode. Place the probe (transducer) at roughly 60 degrees to the carotid artery and observe the time dependence of the blood flow in the artery (listen to the sound corresponding to the pulsation in the loudspeakers).

MEDICAL APPLICATION OF THE ULTRASOUND: THERAPY

Find this part in the chapter 18. SINE WAVE OSCILLATOR.[image: image17.jpg]


[image: image18.jpg]



�





A. Denier's ultrasonoscopic apparatus from 1946.








� echo


� visszhang, echo


� Echo





� transducer


� transzducer


� Transducer





� reflection


� visszaverődés


� Reflexion





� piezoelectric effect


� piezoelektromos hatás


� piezoelektrischer Effekt





�





Fig 10. Colored Doppler-image of the valve of the heart combined with a 2D B-image (left). The frequency shift (velocity) as the function of time (right) in the selected direction (indicated by the blue line).





SUMMARY:





ULTRASOUND: is a mechanical oscillation, or wave that has higher frequency (pitch), than the upper limit of the audible range of the human ear, which is 20 000 Hz. 





DIRECT PIEZOELECTRIC EFFECT: Piezoelectric insulators accumulate electric charge and get polarized when mechanical strain (pressure, tension) is applied on them. This results in measurable potential difference between electrodes placed on the opposite sides of the crystal.  





INVERSE PIEZOELECTRIC EFFECT:  If an electric voltage is applied on the electrodes of the piezoelectric material, it becomes deformed (shrank or stretched).





TRANSDUCER: Here, it is a device that converts electric signal into mechanical, and vice versa.


DAMPING: is a parameter, that characterizes the attenuation of the ultrasound. It is defined as � EMBED Equation.3  ���, where J0 is the incident intensity and J is the intensity after passing through the layer.





ACOUSTIC IMPEDANCE: Acoustic impedance is an important parameter that characterizes acoustic properties of the medium. It is defined as the product of the density (ρ) of the medium, and the velocity (c) of the ultrasound in the medium (Z = ρ · c). The surface of a medium is called boundary surface (or interface) if the acoustic impedance changes.


 


REFLEXIVITY: is a measure of reflection on the boundary surface of media of different acoustic impedances. Reflexivity is defined as the ratio of the reflected and incident intensity as: � EMBED Equation.3  ���, where Z1 and Z2 are the known acoustic impedances of the media (R = 0 for no reflection, and R = 1 for total reflection). 





PULSE ECHO METHOD: The distance (d ) between the reflecting boundary surface and the transducer can be calculated from the elapsed time (Δt) between the emission time of the wave packet-like ultrasound pulse into the medium, and the time of the arrival of ultrasound echo signal, if the velocity of the sound (c) is known in the given medium (Fig 3) as follows: � EMBED Equation.3  ���





DOPPLER EFFECT: is a frequency (and wavelength) change of the wave as a result of the relative motion of the source and the observer.






































�





Fig 8. The reconstructed 3D ultra-sound image of the fetus compared to the photograph of the newborn baby.





�





Fig 7. The M-image (left) represents the movement of the valve in the heart in time at the selected direction, which is indicated in the 2D B image (right).





�





Taking ultrasound images of the human body's soft tissues at 1954 (Life magazine).








�





Fig 14. Top-view of the water-filled phantom.








Ultrasound is a sound that has higher frequency than the upper limit of the hearing range of an average human ear, that is higher than 20 000 Hz (see chapter 22. AUDIOMERTY).








All solids, like piezoelectric materials, have their own resonance frequency (eigen-frequency), which is the frequency of largest amplitude vibration the piece is capable of, that depends on its mass, its dimensions and its elasticity (for example: tuning fork). In order to achieve the highest efficiency of the ultrasound production and detection, the piezoelectric material is excited with an electric signal that corresponds  to its resonance frequency. 








�





Fig 6. The eye of a goat, ultrasound A-image along the optical axis (in the middle), 2D B-image (above) and a one dimensional B-image along the axis of the eye (below). The main echo signals come from the cornea, the anterior chamber, the posterior chamber and the retina.





The higher is the acoustic impedance difference between the two media, the higher is the intensity of the reflected ultrasound. On the liquid–gas or solid–gas interface the ultrasound is totally reflected. Because of this, a coupling medium (gel, water or oil) is used in medical ultrasonography, that helps to transfer the acoustic energy from the transducer to the body and back again. 








�





Fig 15. The image of the back-wall reflection appears in different distances, depending on the material in the finger of the rubber gloves.











Further readings: Damjanovich-Fidy-Szőlősi:


II / 2.4., VII / 1. 2. 4., VII / 1. 7., VIII. / 4.2., 


IX / 5. 1.





�





Fig 11. Doppler flow meter.





�





Fig 1. Reflection and refraction 


of the ultrasound wave (c1 > c2). 


Law of refraction is valid for sound waves as well.








�


Fig 4. In this typical application ultrasound transducer provides the spatial information of the fetus.





Transducer is a signal converting device, which converts the electric signal into the ultrasound and backwards. For example, when the ultrasound is produced, the electric signal in the transducer is converted into the ultrasound. When detected, the ultrasound echo signal is converted to the electric signal in the same transducer.








Direction of propagation of the longitudinal wave is the same as the direction of the oscillation. Sound wave in the air or in the water is longitudinal (but waves on the water surface are not longitudinal!).





�





Direction of propagation of the transverse wave is perpendicular to the direction of the oscillation. Vibration of a string is transverse. 








PAGE  
18. a.  7
Medical physics lab                                                              18. a. ULTRASOUND

[image: image41.jpg]transducer

US beam

water

dwater
water

K

contours of the
rubber glove finger
on the screen
water

dglycerol
glycerol

water A

;'. daicohol
- alcohol



[image: image42.jpg]transducer

£ red

: blood

velocity of the @ cells
blood flow /| fy /

V [ @o®e -\:‘-
0y et "®

blood vessel




[image: image43.jpg]Ci Co

boundary

Cli e
Wa'/e
e refracted Wave w.,

reﬁ\e‘f‘edﬂ




[image: image44.jpg]transducer

ultrasound



[image: image45.jpg]LONGITUDINAL

' wavelength | E direction of
J, o

direction of :
oscillation propagation

| TRANSVERSE



_1361203811.unknown

_1361359285.unknown

_1362230886.unknown

_1362308298.unknown

_1361360052.unknown

_1361358433.unknown

_1361358953.unknown

_1361289033.unknown

_1361200852.unknown

_1361202964.unknown

_1361189886.unknown

